## Slide 1
Winter Workshop PTTF – Marianska 22. – 26. 1. 2018
Miroslav Šos
Diploma thesis
Supervisor : Ing. Petr Böhm Ph.D.

## Slide 2
1/19
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
Development and implementation of data processing for TS in Python
  Temporal fitting of the fast signal
  Incorporation of results from calibration
  Determination of electron temperature and density
  Proper estimation of errors and saving into database
Check TS calibration routines – alternatively perform newly developed methods

## Slide 3
2/19
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
Brief theory of Thomson scattering (TS)
Realization of TS on the COMPASS tokamak
Description of data processing algorithm
Implemented methods for optimisation
Plans and goals for the future
Summary

## Slide 4
Thomson scattering is the elastic scattering of electromagnetic wave on electrons
Due to Doppler effect radiated spectrum is broadened and shifted towards shorter wavelength
3/19
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018

## Slide 5
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
4/19
Thomson scattering on the COMPASS tokamak – overview [Bohm, 2017]
=1064nm
E = 1.5 J
f = 30 Hz
σ = 7 ns
Notes: 100 m vysoko, metru za sekundu.

## Slide 6
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
5/19
Duplexing – two spatial point led by different length of fibres (±13 m ≈ 66 ns)
TS Diagnostics scheme on COMPASS tokamak [COMPASS web, 2014]
1 GHz sampling
Notes: 100 m vysoko, metru za sekundu.

## Slide 7
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
6/19
Each polychromator consists of 5 (4) spectral filters, which spectrally divide the signal
Spectrum of the scattered radiation can be reconstructed
Spectral transmisivity of one polychromator
Notes: 100 m vysoko, metru za sekundu.

## Slide 8
7/19
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
Notes: 100 m vysoko, metru za sekundu.

## Slide 9
8/19
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
TS raw data – 1 GHz sampling
Lasershot time axis
  usually 16 laser trigger marks for each laser
Currently in IDL → transfer to Python
Notes: 100 m vysoko, metru za sekundu.

## Slide 10
2/11
Miroslav Šos: Optimalization of operation of the TS diagnostics / RS defense / 18.9.2017
Notes: 100 m vysoko, metru za sekundu.

## Slide 11
2/11
Miroslav Šos: Optimalization of operation of the TS diagnostics / RS defense / 18.9.2017
To be implemented
Notes: 100 m vysoko, metru za sekundu.

## Slide 12
10/19
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
Purpose of optimisation:
  Correction of any effects leading to incorrectly determined results
  Emphasis on modularity and variability of the system
  Higher requirements on robustness during fitting procedures
  If possible, simplify the code and make it more clear

## Slide 13
11/19
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
Until now fitting by single Gaussian function
Double Gaussian implemented
Originally mitigation
of the stray light effect
Notes: 100 m vysoko, metru za sekundu.

## Slide 14
11/19
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
Until now fitting by single Gaussian function
Double Gaussian implemented
Originally mitigation
of the stray light effect
Not stray light
→ feature of detector
[LAPD 2017]
Notes: 100 m vysoko, metru za sekundu.

## Slide 15
12/19
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
Stray light mitigation → laser misalignment
Distance (temporal) between useful signal and stray light ≈ 20 – 21 ns => 6 m
Notes: 100 m vysoko, metru za sekundu.

## Slide 16
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
13/19
Inconsistency with interferometric measurements
  => corrections on plasma shape (elongation) has been implemented
Notes: 100 m vysoko, metru za sekundu.

## Slide 17
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
14/19
Red color = region without correction
Black color = region with correction
without correction
with correction
Notes: 100 m vysoko, metru za sekundu.

## Slide 18
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
14/19
Works fine when for L-mode with standard plasma density
Notes: 100 m vysoko, metru za sekundu.

## Slide 19
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
15/19
Notes: 100 m vysoko, metru za sekundu.

## Slide 20
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
16/19
For the case of H-mode → whole pedestal becomes cut off
Strong impact of EFIT calculation
Spatial calibration control and revision
Notes: 100 m vysoko, metru za sekundu.

## Slide 21
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
17/19
Further directions of optimisation process:
Mitigation of gaps or jumps in electron temperature and density profiles
→ Robust fitting during spectrum reconstruction
→ Saturated or non-functional channels remove from calculation
→ Revision and verification of spectral calibration
Notes: 100 m vysoko, metru za sekundu.

## Slide 22
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
18/19
Further directions of optimisation process:
Mitigation of gaps or jumps in electron temperature and density profiles
Elimination of configuration within the code and using redundant files
Notes: 100 m vysoko, metru za sekundu.

## Slide 23
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
18/19
Further directions of optimisation process:
Mitigation of gaps or jumps in electron temperature and density profiles
Elimination of configuration within the code and using redundant files
Correct provisional solutions and non-optimal patches
Notes: 100 m vysoko, metru za sekundu.

## Slide 24
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
18/19
Further directions of optimisation process:
Mitigation of gaps or jumps in electron temperature and density profiles
Elimination of configuration within the code and using redundant files
Correct provisional solutions and non-optimal patches
Aim to create simplified and clearly arranged data proccesing system
Notes: 100 m vysoko, metru za sekundu.

## Slide 25
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
Thomson scattering theory and realization of the diagnostics described
Data acquisition process shown
  → Structure and content of data
Data processing scheme
Some methods of optimisation presented
  → Double Gaussian fitting
  → Line-averaged density
Plans and goals for the future
19/19

## Slide 26
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
Thomson scattering theory and realization of the diagnostics described
Data acquisition process shown
  → Structure and content of data
Data processing scheme
Some methods of optimisation presented
  → Double Gaussian fitting
  → Line-averaged density
Plans and goals for the future
19/19
Thank you for your attention

## Slide 27
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
 | Core TS | Edge TS
Field of view | -17 – 211 mm | 215 – 322 mm
Spatial points | 24 | 30
Resolution | 9 – 12 mm | 3.5 – 4.0 mm
Thomson scattering on the COMPASS tokamak – field of view [Bohm, 2017]
Notes: 100 m vysoko, metru za sekundu.

## Slide 28
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
Notes: 100 m vysoko, metru za sekundu.

## Slide 29
Miroslav Šos: Optimisation of processing raw data of the TS diagnostics / Marianska / 22. – 26. 1. 2018
Raman calibration provides tool for absolute calibration
  Raman scattering in proper gas (Nitrogen)
  Fast signal is similar to TS signal
→ Analysis if fitted parameters from Raman could be used in TS fitting?
  Position and width
Laser dependent
Channel
dependent
Notes: 100 m vysoko, metru za sekundu.